## Slide 1: Gluco High Level Tech Foundation
Last Revised: 10/15/24

## Slide 2: Goal: Non-invasive Glucose Monitoring
3 main technological pieces:
Software (app & data acquisition)
  Gluco app, data entry from other CGM and health apps (Apple Health)
Hardware (data acquisition)
  Spectroscopy data from nsp23m (PSD, absorption spectra data)
Machine Learning (data analysis)
  Data pre-processing and ML for estimation/prediction
  Uses both SW add HW data for training, testing & evaluation
Key Takeaway: Software and Hardware are both means of data acquisition, both data sets will be used to train ML model to predict glucose. There is currently no dataset that pairs glucose with spectral data on a large scale.

## Slide 3: Software (Gluco App)
Data Collection:
Ingests CGM data (and other data) from Apple Health, direct user input, or Dexcom's API
Ingests spectral data from the Gluco device
Structures incoming data into a format suitable for analysis
  CGM data has higher status as authority of truth (assuming CGM devices are FDA approved)
  Other related data (Apple Health, HRV, etc.)
Sends data to cloud database for further data analysis, visualization, and ML model development
(Lower Priority) Provides users insights on data captured
Data Usage:
Used in ML training, testing and evaluation.
Used in reports & trends for past.
Provide user insights and trends on existing data
Possible LLM in app for general* glucose and health related questions
Key Takeaway: Software captures CGM data (among other types) from a multitude of sources, captures spectral data from the gluco devices, structures data for further analysis and ML model development, and provides users with insights.

## Slide 4: Hardware (Gluco Device)
Data Collection:
Spectroscopy, nsp32m, used for chemical analysis by spectrophotometry.
  Measure contents of interstitial fluid (ISF) using near infrared and infrared LED’s and nsp32m
  Glucose absorptivity or molar attenuation coefficients:
    0.0002 at 660 nm
    0.0001 at 940 nm
  Beer-Lambert Law - chemical analysis measurements to determine concentration of chemical species that absorbs light
Where epsilon = absorptivity (values above)
L = sample length ( cm)
C = concentration
I_0 = Light transmitted through pure solvent
I = Light traveled through sample
Key Takeaway: Leveraging existing field of chemical analysis by spectrophotometry to measure concentration of glucose in ISF.

## Slide 5: Potential Hardware Sources of Error
Beer-Lambert Law is used for homogenous solutions, ISF and blood are not homogenous although regulated.
We need to develop a calibration curve for Beer-Lambert Law to apply*
Key Takeaway: Empirical methods may not be successful due to their limited applications in heterogeneous solutions. Still we are mainly using this data for ML, not empirical analysis.

## Slide 6: Machine Learning
Mainly focused on prediction algorithms
Using both R and Python
  R for data visualization and statistical analysis
  Python for production and efficiency. Will contain main algorithm
Data coming from software and hardware data collection
  Data used for training, testing and evaluation
Starting ML algorithms with data from Kaggle Competition
  Glucose estimation, from recorded glucose data
  Capture and leverage learnings to build preliminary ML algorithms
Key Takeaway: Use prediction algorithms trained with SW and HW data. Before SW and HW data acquisition, build preliminary ML algo from Kaggle dataset.